## Sheet 1
Question 1: | What is the Make and Model of your XRF Machine and the software name that goes with the machine? | Do you have a service contract for the machine/ remote access (webX capabilities)? | Do you have any issues with breakdowns, customer service/tech support for the machine? | Did you receive training on the new equipment? | What is your sample preparation method? | Do you run any other test methods with your XRF? | How many cement, ash, slag samples do you run annually? | Contact Info
Alabama | ARL PERFORM X 2500 XRF UNIQUANTOMETER; SOFTWARE: XR-OXSAS (runs on Windows 7 Professional 64 bit; SEE QUOTE FOR DETAILS!) | Not yet for this new one, but we will (had one for the older one), but did not/do not have remote access capability for either. | None so far; very few with older one and service tech was excellent as far as response times and proficiency. | Yes. Instrument representative spent a week here at ALDOT providing hands-on training and questions/answer sessions. Also helped reset curves in the new system. | Samples are prepared as pressed pellets and/or glass beads. | Yes. We analyze liquid asphalt samples for PPA content, we chemically analyze steel, copper, fly ash, for oxide content, diesel fuel for ULS content, and can analyze glass beads for heavy metals (unofficially—method not yet approved federally). | 1300-1500 yearly of cement and fly ash. | Kirk Casey 334-206-2349 caseyk@dot.state.al.us
California
Colorado
Florida | Panalytical PW2400, SuperQ software | Yes | Some breakdowns due to age, but customer support is excellent. | Yes | Pressed pellet | We have additional software for the XRF that allows us to test various materials such as soil, glass beads, or any unknown sample. | Approximately 1500 | Awilda Merced awilda.merced@dot.state.fl.us 352-955-6663
Georgia
Idaho | We have an Oxford Twin-X that has a built-in computer running on Oxford software. | We placed a 1-year service contract on the instrument after the last breakdown. It is not set up for remote access. | The hardware for the computer fried when it was 8 years old. Tech support has been OK for the most part. We have had only a couple of breakdowns over the past 20 years with Oxford. They make a good product. | No, but the instrument came with a test method already set up to run cement. They do offer a limited amount of training but the equipment is easy to use. More training would be nice. | Pressed pellet method using an Oxford binding agent. | We have used the XRF to test TiO2 levels in paint. We have also used the XRF to test PO4 levels in binder samples. | Roughly 100 | Eric Dahlinger, Chemist 208-334-8245 Eric.Dahlinger@itd.idaho.gov
Illinois | Oxford Instruments, MDX-1080 Multi-Dispersive (both WDXRF and EDXRF for heavy metals); XpertEase software. | No service contract. Has remote access via SYMANTEX PC ANYWHERE. | Breakdown issues: a. So many interlocks for safety, when one doesn’t work, the whole machine will not work. b. Specific WDXRF elemental detector windows are very thin and sensitive (easily punctured). c. On-site engineer service visits: 5 times since purchase in 2006. d. Service visits for: no P-10 (argon w/ 10% methane) gas flow due to punctured detector window; replacement of P-10 gas flow PC board; replacement of High Voltage detector assembly due to drifting; replacement of vacuum sensor and driver board; preventative maintenance inspection; replacement of power supply assembly; vacuum leak. Service engineers are very helpful and willing to give free phone support. Site service visits are $5k-15k. | Yes, on-site training for two users. | Fused bead for cement/fly ash; pressed pellets for other materials. | Yes, stainless steel and low alloy steel: % nickel, chromium and molyddenum; asphalt: toxic heavy metals, % additives (like used motor oil); latex traffic paint: toxic heavy metals in white or yellow pigments and in the traffic paint film; thermoplastic traffic paint: toxic heavy metals, % TiO2; glass beads used in traffic paint: toxic heavy metals; rock salt: toxic heavy metals. | Cement: 100-150; fly ash: 200-270. | Ronald Price, Field Inspection Coordinator, Ronald.Price@illinois.gov 217-782-7961. Violet Goodman, Instrument Laboratory Chemist, Violet.Goodman@illinois.gov 217-782-7215.
Illinois Tollway
Indiana
Iowa | The XRF is manufactured by PANalytical, it is model type Axiosmax. A recent model and purchase for the lab, I believe 2012. The software is called SuperQ | We do have a service contract. I was not here when the purchase was made so I do not have details. Our service contract includes technical assistance via phone, and if needed a visit from a technician. In our contract we also receive two preventative maintenance visits by a PANalytical technician per year. Remote access (webX) is available with PANalytical and we have used it. | I would say the machine has performed very well thus far. I have been with the department since 2013 and find it works well. However it is nearly brand new. I might mention that we purchased a fusion machine during the purchase of the XRF. The fusion instrument is also made by PANalytical and is a model Eagon2. It is an “automated” fusion machine, you place the crucible and mold into holders and the samples are imported to the furnace. Wetting pellets are added by a hole on the top of the instrument. The pellet sits in wait for the programmed time to be added. There was a cam that would move to drop it. In less than two years the Eagon 2 had a major corrosion of the cams and was not functional! Unfortunately, the fusion instrument was on a separate service contract that expired after one year. When contacting PANalytical about the issue they were very helpful and admitted there was a burgeoning issue being discovered with the design. They sent in a tech and replaced the cams with a different functioning part (paddles) that are more resistant to corrosion. They sent techs in from the east coast and did not charge us for this upgrade. PANalytical also attended to the matter in a short time span. I would say in regards to this particular issue I was very pleased with service. However, I would again caution that our machine is relatively new and we have not had many opportunities to test their customer service. | I was not here during this time, I do think we received support in setting up the instrument and remote support for the calibration. I am unsure how much hands on training took place. I would think there was training and assume you could request this in a contract. | Briefly, our preparation method is to weight out 5.6000 g flux (66% LiTet/34%LiMet)and add 0.8000 grams of a sample that has gone through LOI. For daily samples we add these together in a stoppered glass vial and mix. Samples are then fluxed (1120 C for cement and 1200 for fly ash). Would you like the times for melt, mix, pellet addition, and cooling? Let me know how much detail you may like? | Cement and fly ash are main materials. Our geologist prepares powder samples of limestone and dolomite. We have also informally used the instrument for gaining the chemistry of fine coatings on quartzite aggregates. We will calibrate the instrument for slag and wood in the near future. If you are so bold, the XRF model can also take on liquids. | We run approximately 75 cement samples and between 100- 150 fly ash samples. The geology section runs over 1,000 powdered samples per year. | Michelle Barger, Ph.D. 515-239-1163 michelle.barger@dot.iowa.gov
Kansas | We have a Thermo ARL Advant’XP. Wave length dispersive XRF with a vacuum polychrometer. The software is the old and out dated WinXRF. | Yes, we do have a full service contract with Thermo. | It breaks down all the time. Our old Phillips was the same way. Count on breakdowns…that’s my advice. | Yes, we paid for and received training in Flint MI; It was for the most part, as waste of time and money. | We do fused beads for Portland cement and pozzolans, pressed pellets for CCA treated wood. | No | ~200 to 250 | Joe Beilman Beilman@ksdot.org
Louisiana | PANalytical Magix 2403 (with sample changer and Liebert UPS) | Annual service contract with webX capabilities | No…excellent customer support (both mechanical and applications) | Yes, training at their facility for two employees. | Fused bead (PANalytical PerlX3 Bead Fusion instrument) | No | Fiscal Year 2015 Aggregates 262 Cement 581 Fly Ash 367 Slag 43 Lime 47 Microsilica 18 Thermoplastic 100 | Michael Bailey 225-248-4201 Michael.Bailey@LA.GOV
Michigan | Does not have an XRF Machine.
Minnesota | Bruker S4 Explorer - SPECTRAplus | Yes, we have a Premium Contract with remote (Webex) diagnostic capabilities. | There has been occasional repairs but service has been very good. | A software training class was included with the purchase of the instrument. A separate XRF class was purchased and given by ICDD. | Fusions for cement and fly ash. Loose powder, solid, or pressed pellet for other samples. | Yes, stainless steel, aluminum, low alloy steel, PPA in asphalt binder, semi-quantitative for unknowns, and potentially REOB in asphalt binder. | Approximately 400. | Allen Gallistel Chemical Lab Director allen.gallistel@state.mn.us 651-366-5545
Missouri | Does not use XRF for cement/fly ash analysis. Currently use ICP-OES emission spectrometer for rapid analysis on cements, etc.
Montana
Nebraska | Bruker S4 Pioneer – software is SpectraPlus. | Yes | We occasionally have had to have parts fixed/replaced. Bruker has been excellent at getting a technician on site, usually, within the next 2 days. | Onsite training was conducted during the installation. Further training conducted in Madison, Wisconsin was included in the purchase price, but we didn’t have the budget to travel to that location. | Currently, we use pressed pellet for our cements and fly ashes. We are budgeting to switch to fused beads. | We analyze TiO2 in our traffic paints by ASTM D49764, which is a dry film. We also analyze hydrated lime and crushed stone by powder analysis. | Cements, including Type I, I/II, III, and blended, are 400 – 550 annually. Ashes are less than 50 annually. We only analyze ash used in soil stabilization. Slags are not analyzed. | Jasmine Dondlinger Nebraska Dept. of Roads Materials & Research Division Highway Chemical Test Mgr. 402-479-4874 jasmine.dondlinger@nebraska.gov
Nevada
New York
North Carolina | Thermo Electron AdvantX+ (wavelength dispersive X-ray); Oxsas 2.1 v2.1.5.2432A (new upgraded user software); Uniquant 5.0 software (add-on for $20k) | We have an expensive service contract (quick response type). NCDOT IT doesn’t allow the internet connection on equipment so simple repairs need to have a site visit. | PM’s are done 2x/yr. Occasional breakdowns occur but response to repair it is fast. We have had very good service since we purchased it in 2005. | Yes, 2 weeks for XRF. One week for fluxer equipment. This was included in the bid price and was spec’d out. | Specifically for cement, lithium tetraborate and cement (special ratio) is used to cast “beads”, 32 mm diameter, in a fluxer (Claisse brand) using platinum ware. We can press out the cement powder, but that’s for a backup method only. We now have two fluxers so pressing is out of the picture. Casting beads removes matrix effects (elemental interference) and creates nice calibration curves using a combination of NIST standards and old CCRL cements. | Yes. Paint pigments, metals in asphalt liquid binder, zinc dust, metal checks and unknowns on the Uniquant 5.0 software. In addition to cement, we analyze quick lime, chlorides in concrete, fly ash, grout, and cement content in concrete using calibration curve methods. | Cement- 1200, Fly Ash/Slag- 350, Chloride in Concrete- 475 | Sharon Freeman 919-329-4090 sdfreeman@ncdot.gov
North Dakota | Thermo Scientific--ARL Advant’X with Oxsas and UniQuant software programs. | We do have a service contract and they (technicians) are able to access the XRF remotely. | The instrument does need to be recalibrated a little more often than expected. However, the technical support has been above average, or better, and no major breakdowns to report. | There was training on the new instrument, however, more would be beneficial. There are many intricacies with software. | Pressed pellet with cellulose binder. | Cement and UniQuant for paint chips along a few other random samples. | About 50-70 per year | Ryan Johnson 701-328-6923 rejohnson@nc.gov
Ohio | Bruker S4 Explorer - SPECTRAplus | No | Have had issues throughout performance life . Support has been excellent for the most part. | Yes. | Fused beads and pressed powder pellets. | Yes, aluminum, NaCl, SO3 in soil, cement, glass beads, Unknowns. | Approximately 350. | Maria Kerestly, Chemical Lab 614-275-1349 maria.kerestly@dot.state.oh.us
Oklahoma | Panalytical Axios DY1422 wavelength dispersive XRF, superQ & protrace software | Yes, we have remote access capabilities on demand our instrument is not constantly connected to internet. | Yes we have issues with instrument, customer service/tech support no issues | Yes | Fused glass bead and pressed pellet | Yes, glass beads fly ash, & unknown samples | Approximately 200 | Danna Crouse Environmental Chemical Lab Scientist IV 405-522-4915 dcrouse@odot.org
Pennsylvania | PANalytical Axios, Super Q, Version 4.0 with IQ+ and Pro-Trace. | Service is done on an as needed basis, web X is not permitted due to network security. | PANalytical has always provided excellent support for both service and tech. | Yes, the unit was setup here by an applications specialist who provided training and assisted in initial calibrations. | Our primary and preferred method is fused bead. We also have the capability to make pressed powder pellets. | We have also developed methods for whole rock, steel, aluminum, polymers, dry film paints, treated wood, ETC. Actively working on development methods for analysis of recycled engine oil bottoms in asphalt binders by ashing. | Over 900. | Winson King, Chemist III 717-783-6716 wiking@pa.gov
Rhode Island
South Dakota | Does not perform cement chemistry testing.
Tennessee | ARL ADVANT’XP | Yes | No | Yes | Pressed Pellet | No | Cement (496), Fly Ash (384), Slag (5), Limestone (94) | Bobby Anderson 615-350-4181
Texas | We have two: a Rigaku Primus II XRF(a tube above 4kW system) the software is called ZSX and a Bruker S4 Explorer XRF(a tube below 1kW system) the software is called SpectraPlus. Tube above systems keeps things from falling into the tube and other parts if it falls apart. I have only had that happen once but it was a huge hassle to clean up. Both Bruker and Rigaku manufacture good systems, either one would be a great addition. Having used both it seems the Rigaku software is less “buggy”. Creating a calibration curve seems less difficult with the Rigaku software but is less intuitive to making changes after the calibration is complete. | We do not have a service contract on the Rigaku instrument as it is only about 1.5 yrs old. We do have a “tubes for life” contract for the Bruker(the instrument is about 8 yrs old) that will replace the tube if it goes out for the next 1-2 yrs. We have had to replace the tube a couple of years ago in the instrument and seemed like a reasonable assumption that we would need to replace the tube again at some point. Right now, I’m hoping the tube does go out so we can get at least one more new tube out of the contract. We do have WebX capabilities for both instruments. I have used Webex on our Bruker XRF multiple times and it was an invaluable help. | We have had several issues with the Bruker instrument but they have always been able to either help over the phone or webex or were able to send someone out to repair the instrument. I think in the 8 or so years that we have had the Bruker they have been out 3 -4 times to service it. I have had to get Rigaku out once to replace a bad vacuum pump. The instrument was still under warranty. Customer Service for Rigaku has been excellent. I have contacted them a couple of times with software questions and also application questions and received responses quickly that were helpful. | We purchased training with both instruments. I think both provided excellent training on the use of their software and also XRF in general. | Sample prep for Portland cement, blended cement, and fly ash is done by Lithium borate fusion. Slag we press a pellet. | Between the two instruments we test the alloy of aluminum sign blanks, TiO2 content of traffic pavement marking material, check for PPA and REOB in asphalts, and the chloride content of grout samples. | We run about 950 cement, fly ash and slag samples per year. | Cliff Coward 512-506-5855 Cliff.Coward@txdot.gov
Utah
Washington
Wisconsin